18英语1 20181507105 刘佳慧

原文：

Response:

GVK SOP BR030 was updated on July 20, 2012 to version 08, a few months prior to the H052-12 and H053-12 BE studies, to address the Agency’s observation on documentation of acquisition errors. We inquired GVK about it. GVK indicated that acquisition errors in terms of sequence stoppage were documented per SOP BR030. All records of sequence stoppage during the bioanalyses of H052-12 and H053-12 studies are provided as follows:

Sequence stoppage and re-submission for H052-12, the fasting study in Attachement-1:

- Run 1 on LC/MS/MS instrument 4A was stopped between Injection 84 and 85, probably due to the communication issue between HPLC and LC/MS/MS, so the system was restarted and the sequence resubmitted from Injection 85.

- Before the full sequence run of Run 21 on LC/MS/MS instrument 4A, there were stoppage twice, the first time was due to unknown software error and the second time was because the purge valve was inadvertently open. The third submission of the full sequence run of Run 21 was successful and data submitted in the ANDA.

Sequence stoppage and re-submission for H053-12, the fasting study in Attachement-2:

- Run 32 on LC/MS/MS instrument 3 was stopped between Injection 72 and 73 due to significant back pressure of column, so the system was restarted and the sequence resubmitted from Injection 73 after changing the peak tubing’s of the HPLC.

- Run 19 on LC/MS/MS instrument 3 was stopped between Injection 58 and 59: During the acquisition of Run19 the auto sampler tray was inadvertently removed by the analyst with an intention of loading the Run 20 samples in the second tray. This was identified when the sequence was stopped automatically with acquisition error, therefore the sequence was resubmitted from Injection 59.

- Before the full sequence run of Run 4 on LC/MS/MS instrument 3, there were errors in input of the rack position, plate code and plate position. So the full sequence was stopped re-submitted from Injection 1 again.

- Before the full sequence run of Run 7 on LC/MS/MS instrument 3, carryover of AQSTD was observed due to air filter blockage. The full sequence run was resubmitted after sonication of the air filters.

In the mean time, Prinston went through the 100% chromatograms of H052-12 and H053-12 studies and the accompanying records, identified some potential concerns whereby the sequence run didn’t seem to be conducted or processed perfectly, and asked GVK for explanation or clarification as indicated in Table 5-1 and Table 5-2 above. GVK did respond these specific questions to our satisfactory as follows:

The potential concerns in data acquisition for H052-12, the fasting study:

- Why was there data re-processing for Valsartan in Run 22 for Subject 52&53: During data review, the analyst found that the Period 1 at 4.5 hour sample of Subject No. 53 was not integrated for Valsartan due to slight shift in retention times of drug and internal standard. The shift in the retention time is within the limits per the Method SOP, so considered as acceptable. To integrate this sample the integration parameters were modified and the entire batch was reintegrated for consistency across all samples. Since the calculation of drug concentration in a given sample depends on the peak area ratio of drug over internal standard, and slight shift in retention time for both drug and internal standard was observed, this data re-processing will not affect the passing BE results of fasting study H052-12.

- Why Run 28 for repeated analyses of Subject 02&03 failed Valsartan calibration curve: After an investigation, GVK thinks that the failure in Valsartan calibration curve was probably due to some sample preparation errors which could occur during pipetting of plasma sample into vial, addition of internal standard, addition of extraction solvent, vortexing, transfer of supernatant, evaporation and reconstitution of the residue. The failure was not due to analytical errors such as poor chromatography and or internal standard variation. The possible sample preparation error didn’t occur for HCTZ analyses, per SOP BR016 for Project Sample Analysis, Data Checkup and Acceptance Criteria, the data generated for HCTZ was accepted, whereas Valsartan was to be re- analyzed for subject 02&03.

- 译文：

- 答复：

- 2012年7月20日，即进行H052-12和H053-12 BE研究的几个月前，GVK的SOP BR030更新至第8版，以解决机构对采购错误相关文件的观察。我们向GVK询问了此事。GVK表示，它们按照SOP BR030记录了序列停止方面的采集错误。H052-12和H053-12研究的生物分析期间的所有序列停止记录提供如下：

- 附件1中，H052-12空腹研究中的序列停止和重新提交情况如下：

- 在4A号液相色谱质谱联用仪 （LC/MS/MS instrument）上的第一轮运行在第84次注射和第85次注射之间停止，可能是由于液相色谱（HPLC） 和联用仪之间的通信问题，因此系统重启并从第85次注射开始重新提交序列。

- 在进行4A号液相色谱质谱联用仪的第21轮全序列运行前，有两次序列停止，第一次是由于未知的软件错误，第二次是因为吹扫阀意外打开。 第21轮全序列运行的第三次提交成功，数据在简化新药申请 （ANDA）中提交。

- 附件2中，H053-12空腹研究中的序列停止和重新提交情况如下：

- 由于色谱柱背压较大，3号液相色谱质谱联用仪上的第32轮运行在第72次和第73次注射之间停止，因此系统重新启动，并在改变液相色谱峰管后从第73次注射重新提交序列。

- 3号液相色谱质谱联用仪上的第19轮运行在第58次和第59次之间停止：在第19轮采集期间，为将第20轮的样品装入第二个托盘，分析员无意中移除了自动取样器托盘。但当序列因采集错误而自动停止时，该事物被发现，因此序列从第59次序列重新提交。

- 在3号液相色谱质谱联用仪上的第4轮全序列运行之前，机架位置、板代码和板位置的输入存在错误。因此，整个序列停止，再次从第1次注射重新提交。

- 在3号液相色谱质谱联用仪上的第7轮全序列运行之前，观察到因空气过滤器堵塞造成的AQSTD携带。对空气过滤器进行超声波处理后，重新提交全序列运行。

- 与此同时，普林斯顿浏览了H052-12和H053-12研究的100%色谱图和随附记录，发现了一些潜在的问题，即序列似乎没有完美地运行，并要求GVK进行解释，或如上表5-1和表5-2中所示做出解释或澄清。

- GVK对这些具体问题的答复如下：

- 在H052-12空腹研究中，数据采集的潜在问题：

- 为什么在第22轮中对52号和53号受试者的缬沙坦数据进行重新处理：分析员在数据审查期间发现，由于药物和内标物的保留时间略有变化，53号受试者在4.5小时取样时的第1个周期未纳入缬沙坦。但此变化在SOP的限制范围内，因此被认为是可接受的。而为了整合该样本，还对整合参数进行了修改，并对整个批次进行了重新整合，以确保所有样本的一致性。由于给定样本中药物浓度的计算取决于药物与内标物的峰面积比，并且观察到药物和内标物的保留时间略有变化，因此该数据的重新处理不会影响H052-12的空腹研究结果的通过。

- 为什么在第28轮中对02号和03号受试者的重复分析使缬沙坦校准曲线不合格：经调查，GVK 认为，缬沙坦校准曲线的不合格可能是由于分析员在将血浆样品移液到小瓶中时，除添加内标物、添加萃取溶剂、搅拌、转移上清液、蒸发和重构残留物等因素外，还可能会发生的一些样品制备错误。不合格并不是由于分析错误，如色谱分析不当和（或）内部标准偏差。在进行氢氯噻嗪（HCTZ）分析时，没有出现可能的样品制备错误，而根据项目样品分析、数据检查和验收标准的SOP BR016，该氢氯噻嗪数据被接受，而02和03号受试者的缬沙坦数据则需重新分析。